Power On Initialize (I Steps)

Step	Description	Details/Actions	PTR	AS	PTD	V1	V2	V3	V4	V5	V6	S1	S2	S3	S4	S5	S6	P1	X	Y	Z	A
I1	Power Up	If S2 OFF / Goto I2 / Else Goto I3				OFF	OFF	OFF	OFF	OFF	OFF	OFF	OFF	OFF	OFF	OFF	OFF	<10	?	?	?	?
I2	Raise Needle	Raise Z Until S2 On / Or until Z MaxStep / If S2 On Goto I3 / Else NEEDLE ERROR HOLD				OFF	OFF	OFF	OFF	OFF	OFF	OFF	OFF	OFF	OFF	OFF	OFF	<10	?	?	?	?
I3	Rinse Pressure / Check	If P1 <10 / Alert to Pressurize / OK to continue I4				OFF	OFF	OFF	OFF	OFF	OFF	OFF	ON	OFF	OFF	OFF	OFF	<10	?	?	Up	?
I4	Vial Check	If S1 Off / Vial Not Found / Else Vial Found / Either Go to I5				OFF	OFF	OFF	OFF	OFF	OFF	OFF	ON	OFF	OFF	OFF	OFF	>10	?	?	Up	?
I5	Initialize / Sample / Syringe	If S4 Off Goto I6 / Else SYRINGE ERROR HOLD				OFF	OFF	OFF	OFF	OFF	OFF	OFF	ON	OFF	OFF	OFF	OFF	>10	?	?	Up	?
I6	Home / Sample / Syringe	Lower Y Until S4 On / Or until Y MinStep / If S4 On Goto I7 / Else SYRINGE ERROR HOLD				OFF	OFF	OFF	OFF	OFF	OFF	OFF	ON	OFF	ON	OFF	OFF	>10	?	DN	Up	?
I7	Zero / Sample / Syringe	Raise Y (Syringe) / To Y ZeroStep / Goto I8				OFF	OFF	OFF	OFF	OFF	OFF	OFF	ON	OFF	OFF	OFF	OFF	>10	?	Up	Up	?
I8	Initialize / Calibration / Standard / Syringe	Out: Cval (Volume) / Last Reported / OK to continue I9 / Recal Goto CS1				OFF	OFF	OFF	OFF	OFF	OFF	OFF	ON	OFF	OFF	OFF	OFF	>10	?	Up	Up	?
I9	Power On / Complete	CSSyringe OK, / Goto SB1				OFF	OFF	OFF	OFF	OFF	OFF	OFF	ON	OFF	OFF	OFF	OFF	>10	?	Up	Up	OK