Ready Mode(RD)

Step	Description	Details/Actions	PTR	AS	PTD	V1	V2	V3	V4	V5	V6	S1	S2	S3	S4	S5	S6	P1	X	Y	Z	A
RD1	Vial Check	If S1 Off / Vial Not Found / OK to continue RD2 / Else Vial Will Be Skipped / OK to continue SB2				OFF	OFF	OFF	OFF	OFF	OFF	On/Off	ON	OFF	OFF	OFF	OFF	<10>10	OK	Up	Up	OK
RD2	Rinse Pressure / Check	If P1 <10 / Rinse Pressure Low / Fixed Rinse OK? / Goto SB3 / Else Goto RD3				OFF	OFF	OFF	OFF	OFF	OFF	On/Off	ON	OFF	OFF	OFF	OFF	<10>10	OK	Up	Up	OK
RD3	System / Ready / To Run / Options / Page	Buttons: / Run Goto RD5 / Align Vial Goto AV1 / Fill Cstd Goto CS1 / Sample Volume Goto RD4				OFF	OFF	OFF	OFF	OFF	OFF	On/Off	ON	OFF	OFF	OFF	OFF	>10	OK	Up	Up	OK
RD4	Set Sample / Volume	Set Sample Volume: / 5mL Set to 5+RV / 10mL Set to 10+RV / 25mL Set to 25+RV / Goto RD1 After Selection				OFF	OFF	OFF	OFF	OFF	OFF	On/Off	ON	OFF	OFF	OFF	OFF	>10	OK	Up	Up	OK
RD5	Start Run	Load Vials and / Press Start / Start Goto RS1 / Else Wait				OFF	OFF	OFF	OFF	OFF	OFF	On/Off	ON	OFF	OFF	OFF	OFF	>10	OK	Up	Up	OK