Run Sample Mode(RS) Page 1 of 3

Step	Description	Details/Actions	PTR	AS	PTD	V1	V2	V3	V4	V5	V6	S1	S2	S3	S4	S5	S6	P1	X	Y	Z	A
RS1	Advance X / Vial Chain	If S2 = ON / Advance X to XINC / Goto RS2 / Else ERROR Hold				OFF	OFF	OFF	OFF	OFF	OFF	On/Off	ON	OFF	OFF	OFF	OFF	>10	>>	Up	Up	OK
RS2	Check for / Vial	If S1 = ON / Goto RS3 / Else No Vial Run Ended Goto SB1				OFF	OFF	OFF	OFF	OFF	OFF	On/Off	ON	OFF	OFF	OFF	OFF	>10	OK	Up	Up	OK
RS3	Wait for P&T Ready	If P&T Pin = Ready / Goto RS4 / Else Waiting for P&T Ready Hold	RDY			OFF	OFF	OFF	OFF	OFF	OFF	On/Off	ON	OFF	OFF	OFF	OFF	>10	OK	Up	Up	OK
RS4	Move Needle / Down	Move Z>Down				OFF	OFF	OFF	OFF	OFF	OFF	On	OFF	OFF	OFF	OFF	OFF	>10	OK	Up	DN	OK
RS5	Fill Sample / Syringe	V6 Needle Flush / Move Y>SV+ @Sp=W				OFF	OFF	OFF	OFF	OFF	ON	On	OFF	OFF	OFF	OFF	OFF	>10	OK	SV+	DN	OK
RS6	Check If No Std Ring / Skip Std	If S3 = On / Goto RS9 / Else Goto RS8				OFF	OFF	OFF	OFF	OFF	OFF	On	OFF	On/Off	OFF	OFF	OFF	>10	OK	SV+	DN	OK
RS7	Fill Std. / Loop	V3 Standard Valve / Move A>A+AINC				OFF	OFF	On	OFF	OFF	OFF	On	OFF	Off	OFF	OFF	OFF	>10	OK	SV+	DN	INC
RS8	Start P&T	Send P&T Start		ST		OFF	OFF	On	OFF	OFF	OFF	On	OFF	Off	OFF	OFF	OFF	>10	OK	SV+	DN	INC
RS9	Rinse to Vial	V4 Waste / Move Y<SV @Sp=W				OFF	OFF	OFF	On	OFF	OFF	On	OFF	Off	OFF	OFF	OFF	>10	OK	SV	DN	OK
RS10	Sample / Transfer	V4 Waste / V1 Syringe / Move Y<0V @Sp=W				On	OFF	OFF	On	OFF	OFF	On	OFF	Off	OFF	OFF	OFF	>10	OK	>0	DN	OK

Run Sample Mode(RS) Page 2 of 3

Step	Description	Details/Actions	PTR	AS	PTD	V1	V2	V3	V4	V5	V6	S1	S2	S3	S4	S5	S6	P1	X	Y	Z	A
RS11	Sampler Rinse Gas Fill	V4 Waste / V5 Gas Flush / Move Y>5mL @Sp=G				OFF	OFF	OFF	On	On	OFF	On	OFF	Off	OFF	OFF	OFF	>10	OK	>5	DN	OK
RS12	Sampler / Gas Rinse 1	V4 Waste / V1 Syringe / Move Y<0mL @Sp=G				ON	OFF	OFF	On	OFF	OFF	On	OFF	Off	OFF	OFF	OFF	>10	OK	>0	DN	OK
RS13	Needle Rinse Gas Fill	V5 Gas Flush / Move Y>5mL @Sp=G				OFF	OFF	OFF	OFF	On	OFF	On	OFF	OFF	OFF	OFF	OFF	>10	OK	>5	DN	OK
RS14	Needle / Gas Rinse 1	Move Y<0mL @Sp=G				OFF	OFF	OFF	OFF	OFF	OFF	On	OFF	OFF	OFF	OFF	OFF	>10	OK	>0	DN	OK
RS15	Move Needle For Rinse	Move Z>R Rinse Pos.				On	On	OFF	OFF	OFF	OFF	On	OFF	OFF	OFF	OFF	OFF	>10	OK	>0	>R	OK
RS16	Needle Rinse H2O Fill	V2 Rinse H2O / V1 Syringe / Move Y>5mL @Sp=W				OFF	OFF	OFF	OFF	OFF	OFF	On	OFF	Off	OFF	OFF	OFF	>10	OK	>5	>R	OK
RS17	Needle H2O Rinse	Move Y<0mL @Sp=W				OFF	OFF	OFF	OFF	OFF	OFF	On	OFF	Off	OFF	OFF	OFF	>10	OK	>0	>R	OK
RS18	Needle / Rinse 2 Gas Fill	V5 Gas Flush / Move Y>5mL @Sp=G				OFF	OFF	OFF	OFF	On	OFF	On	OFF	OFF	OFF	OFF	OFF	>10	OK	>5	>R	OK
RS19	Needle / Gas Rinse 2	V6 Needle Flush / Move Y<0mL @Sp=G				OFF	OFF	OFF	OFF	OFF	On	On	OFF	Off	OFF	OFF	OFF	>10	OK	>0	>R	OK
RS20	Raise / Needle Up	Move Z>U Up Pos.				OFF	OFF	OFF	OFF	OFF	OFF	On	ON/ / OFF	Off	OFF	OFF	OFF	>10	OK	>0	UP	OK
RS21	Check / Needle Position	If S2 = ON / Goto RS22 / Else ERROR Hold				OFF	OFF	OFF	OFF	OFF	OFF	On	ON	Off	OFF	OFF	OFF	>10	OK	>0	UP	OK
RS22	Wait for / P&T Desorb	If P&T Pin = Desorb / Goto RS23 / Else Waiting for P&T DESORB Hold			DS	OFF	OFF	OFF	OFF	OFF	OFF	On	ON	Off	OFF	OFF	OFF	>10	OK	>0	UP	OK

Run Sample Mode(RS) Page 3 of 3

Step	Description	Details/Actions	PTR	AS	PTD	V1	V2	V3	V4	V5	V6	S1	S2	S3	S4	S5	S6	P1	X	Y	Z	A
RS23	Sampler Rinse H2O Fill	V2 Rinse H2O / V1 Syringe / Move Y>5mL @Sp=W				ON	ON	OFF	OFF	OFF	OFF	On	ON	Off	OFF	OFF	OFF	>10	OK	>5	>R	OK
RS24	Sampler H2O Rinse	V4 Waste / V1 Syringe / Move Y>5mL @Sp=W				ON	OFF	OFF	ON	OFF	OFF	On	ON	Off	OFF	OFF	OFF	>10	OK	>0	>R	OK
RS25	Sampler Rinse Gas Fill 2	V4 Waste / V5 Gas Flush / Move Y>5mL @Sp=G				OFF	OFF	OFF	On	On	OFF	On	ON	Off	OFF	OFF	OFF	>10	OK	>5	DN	OK
RS26	Sampler / Gas Rinse 2	V4 Waste / V1 Syringe / Move Y<0mL @Sp=G				ON	OFF	OFF	On	OFF	OFF	On	ON	Off	OFF	OFF	OFF	>10	OK	>0	DN	OK
RS27	System H2O Rinse Fill	V2 Rinse H2O / V1 Syringe / Move Y>5mL @Sp=W				ON	ON	OFF	OFF	OFF	OFF	On	ON	Off	OFF	OFF	OFF	>10	OK	>5	>R	OK
RS28	System H2O Rinse	V1 Syringe / Move Y>0mL @Sp=W				ON	OFF	OFF	OFF	OFF	OFF	On	ON	Off	OFF	OFF	OFF	>10	OK	>0	>R	OK
RS29	System / Gas Rinse Fill	V5 Gas Flush / Move Y>5mL @Sp=G				OFF	OFF	OFF	OFF	ON	OFF	On	ON	Off	OFF	OFF	OFF	>10	OK	>5	>R	OK
RS30	System / Gas / Rinse	V1 Syringe / Move Y<0mL @Sp=G				ON	OFF	OFF	OFF	OFF	OFF	On	ON	Off	OFF	OFF	OFF	>10	OK	>0	DN	OK
RS31	Finish Sample	SAMPLE FINISHED / GOTO RS1